# Final Report

# Data Acquisition Box

(An updated version from 2003)

By

# Gene Bender

# DeSmet Jesuit High School

# July 2, 2004

# Contents

Overview	2

A Hybrid LBNL Cosmic Ray Detector	2

The Detectors	6

Changing PMT Voltage	7

Comparator Threshold of the LBNL Counter	9

Wiring Schematic	10

# Overview

During the summers of 2003 and 2004 I had the opportunity to work with cosmic ray detectors.    Using detectors that were constructed over the past two summers, data was collected using different data acquisition methods.  This report summarizes what I have done in terms of building a data acquisition box utilizing a Lawrence Berkley National Laboratory (LBNL) cosmic ray detector circuit board and a simple 4 digit counter mounted together in a box for easy student and teacher use.  While this summer’s data acquisition box is very similar in design to last years, there are a few changes.  Most significantly the reset button is now simpler to use.  This document also summarizes technical information discovered about the LBNL circuit board and our detectors.  Our goal was to utilize the features incorporated into the LBNL circuit board while expanding the possibilities.

# A Hybrid LBNL Cosmic Ray Detector

The LBNL cosmic ray detector consists of two detectors and a data acquisition circuit board.  The detectors and circuit board are mounted in a wooden box.  The counter has the ability to count cosmic rays from one detector or to count cosmic rays that strike two detectors in coincidence.  Other features on the LBNL circuit board include a 3 digit counter and 5 volt outputs for PMTs.  (For complete information on the LBNL cosmic ray detectors go to http://www.lbl.gov/abc/cosmic/.)

The LBNL circuit board has a 3 digit counter.  When counts run greater than 1000, the thousands digit must be counted by hand.  In order to overcome this problem, we added a 4 digit decade counter to our detector providing a millions digit.   The LBNL board and the 4 digit counter were mounted together in a box.  (See Figure 1)

In order to allow many different kinds of experiments to be performed it was decided that that the two detectors needed to be free to move independently from one another.  Experiments such as changing the vertical distance between the detectors or changing the amount of scintillator overlap, or the amount of horizontal distance between the detectors require that the detectors can be easily moved.  The length of wire used to connect the detectors to the DAQ box can be increased as desired.

## Figure 1 – Top view of box

## Figure 2 –  Switches

| Switch | Function |
| --- | --- |
| PMT | This switch allows you to choose which PMT you wish to count. |
|  | 1 1-2 2 | Counter is triggered by PMT #1 only. (Singles mode) Counter is triggered on coincidence mode using both PMT #1 & PMT #2 Counter is triggered by PMT #2 only. (Singles mode) |
| CNT | TIME HLD CONT | Counter is on 60 second count mode. *see note below Hold mode - Counter is not counting. Counter is on continuous timing mode. |
| RST | Resets both counters to zero. |
| BEEP | Turns beeper on or off. |

* Note that once the 60 second count mode has been started the 60 second timer cannot be stopped.  You can switch modes but the timer cannot do another 60 second count until the full 60 seconds has elapsed.

## Figure 3 – Side View

| Port | Function |
| --- | --- |
| PWR | Input for AC/DC converter. |
| PMT #1 | Input for PMT #1 signal. |
| PMT #2 | Input for PMT #2 signal. |

# The Detectors

Each detector consisted of a scintillator plate glued to a light guide.  The light guide was then connected to a photomultiplier tube with a silicon cookie.  The photomultiplier tube was not glued to the light guide so that the photomultiplier can be changed if you wish to do so.  The detectors used originally had 2 inch photomultiplier tubes connected to them but they were changed to 1 inch photomultiplier tubes.

## Figure 4 - Detector

# Changing the PMT Voltage

The high voltage supply used to provide high voltage to the photomultiplier tubes can be varied from 0 – 2000 volts.  Using a screwdriver turn the high voltage adjustment screw to adjust the output voltage.  To determine the high voltage output connect a voltmeter to the low voltage measurement port.  Each volt measured at the low voltage port corresponds to 1000 volts at the high voltage output.  (See Figure 6)  For example if you measure 1.5 volts at the low voltage port the high voltage output is 1500 volts.

## Figure 5 – High Voltage Supply

## Figure 6 – High Voltage Supply

# Comparator Threshold of the LBNL Counter

In order to determine how large of a pulse is needed to trigger the LBNL counter the following experiment was conducted.  Pulses of varying magnitude were connected to the LBNL DAQ board to see if the counter would count them.  All of the pulses were 70ns long.

| Input Signal (mV) | Counter was triggered |
| --- | --- |
| 40 | Yes |
| 20 | Yes |
| 10 | Yes |
| 5 | No |

Table 1

A 10mV signal is necessary to trigger the LBNL counter.

In order to determine how long of a pulse is needed to trigger the LBNL counter the following experiment was conducted.  Pulses of varying length were connected to the board to see if the counter would count them.  All of the pulses were 10 mV signals.

| Input Signal Length (ns) | Counter was triggered |
| --- | --- |
| 70 | Yes |
| 40 | Yes |
| 25 | Yes |
| 10 | Yes |

Table 2

A 10ns pulse will trigger the LBNL counter, but due to limitations of the equipment used we were unable to determine just how short the signal could be and still trigger the counter.  We were unable to measure a signal less than 10ns using or oscilloscope.  This may have been due to the limits of the scope or the function generator.

Our experiment showed that the comparator threshold of the LBNL counter is a 10mV signal that is 10ns long.

# Wiring Schematic

Figure 7 – Wiring Schematic

# Text boxes

10 ns

10 mV

Low voltage measurement here

Power Supply

AC/DC converter

LBNL Counter

4-digit decade counter

12V DC in

12V DC in

High Voltage out to PMTs

High Voltage

Adjustment Screw

Low voltage measurement here

High Voltage

Adjustment Screw